## Cover
Respondent Name:
Mailing Address:
City, State, Zip:
Telephone:
Email:
Request for Proposal (RFP)
RFP No. 07122024
DI/RO Water Purification System Installation & Service
RFP RELEASE DATE: July 12, 2024
PROPOSAL DUE DATE: July 29, 2024
PROPOSAL DUE TIME: 2:30 PM CST
SUBMIT ALL PROPOSALS TO:
            By FedEx, UPS or another private carrier to Physical Location:
      University of Arkansas – Business Services
      UPTW Room 101
      1001 East Sain Street
      Fayetteville, AR 72703

## Requirements
Is your organization willing and able to adhere to all requirements? If not, please explain.
UA Requirement | Answer (select from dropdown) | Notes:
1. Systems shall deliver either Type One or Type Two water as specified by individual departments. Type One and Type Two water is defined by the College of American Pathologists National Committee for Clinical Laboratory Standards.
2. Vendor shall provide all material, labor, supervision, equipment, tools, and incidentals necessary to maintain and/or repair DI/RO Water Purification Systems for the University of Arkansas.
3. Awarded bidder shall provide next-day service for Tank and Filter Exchanges called in prior to scheduled service.
4. Vendor is responsible for providing service and maintenance – including replacement of filters, bulbs, and other consumables – as per attached specifications of lab sites.
5. Vendor is responsible for performing regular preventative maintenance and site visits as agreed upon by the University to ensure the proper and consistent operation of campus water purification systems.
6. The awarded vendor is responsible for inspecting and testing all water system components during site visits, including storage tanks, filters, pumps, softeners, valves, control systems, and pressure vessels.
7. The awarded vendor is required to regularly test water quality for the appropriate indicators/parameters, including pH, conductivity, resistivity, and total dissolved solids (TDS).
8. The successful vendor shall provide invoices to the University for each service request, which shall include the number of units purchased, catalog number, ordering department, and related purchase order numbers
9. Vendor is responsible for developing and maintaining a service log for each site on campus detailing at a minimum all site visits, repairs, preventative maintenance, and testing performed.
10. The awarded vendor shall provide service tickets following the completion of all system repairs and service. Each service ticket must include a unique number, identify the suspected issue(s), and detail steps taken to resolve the issue(s). All service tickets must be presented to appropriate University personnel for signature.
11. Vendor is responsible for cleaning RO membranes when the normalized flow decreases by 10% or the overall pressure drop increases by 15% from the initial acceptance test conditions.
12. All new components/systems installed shall have a parts and labor warranty for a minimum period of one (1) year.
13. Vendor is responsible for providing replacement DI/RO Water Purification System design and installation support as needed.

## Sections A to D
Organization and Ownership | Answer:
1. State your organization's legal name, address, and state of incorporation. If your organization is a subsidiary of a parent firm provide the same information for the parent company.
2. Please state the name, title, phone number, and email address of your organization’s primary contact for this RFP.
3. Please state the name, title, phone number, and email address of your organization’s secondary contact for this RFP.
4. Provide a description of your organization's core activities.
5. Provide a brief statement of qualifications that includes your organization’s size and footprint in AR.
6. How long has your organization been operational?
7. What are the key differentiators that provide your company with a strategic advantage in the marketplace relative to your competitors?
8. Please include any special accreditations including minority- and women-owned business enterprises (MWBEs) or other.
Financial Capability | Answer:
1. Has your organization filed for bankruptcy within the previous seven (7) years?
2. Include audited Statement of Cash Flow, Balance Sheet, and Income Statement from the prior two (2) years. Submit statements and reports from the corporate entity making the Proposal rather than its parent corporation.
Experience and Qualifications | Answer:
1. Provide a list of three references for whom similar services were provided, preferably of public institutions of higher education, including amount of annual sales. Please include the following information for each reference: contact name, title, phone number, and email address.
2. Provide a list of institutions of higher education with which the firm has a signed term contract.
3. Provide your historical relationship with UA (if any) and list up to 3 examples of projects that you have provided to UA or its associated departments.
4. List any services that would need to be outsourced to a 3rd party to fulfill the scope of work.
Technology | Answer:
1. Does your solution integrate with Workday? Please explain.
a. Please describe your organization’s ability to integrate with Workday for billing purposes.
2. How much experience does your organization have working with Workday? Please quantify.
3. Please describe any technology your organization utilizes to support DI/RO Water Purification System operations.
4. Please describe any technology your organization utilizes to manage service requests.
5. Does your organization offer a customer portal? If so, please describe features and functionality of the portal to review service orders, schedule service, and catalog sales information, etc.
6. Please describe your organization’s reporting capabilities.

## Scope Questionnaire
Scope | Answer:
1. Please describe your organization’s approach to completing the scope of work outlined in this RFP.
2. Please explain your organization’s approach to account management. Describe the Service Team responsible for working with the University of Arkansas. Please include resumes and any relevant information.
3. Please detail key roles and responsibilities of the account team responsible for working with the University of Arkansas.
4. Where is the nearest service center/where will services be performed out of?
5. Please list any certifications or training that your technicians are required to possess related to servicing and repairing DI/RO water purification systems.
6. Please detail your organization’s approach to performing required maintenance.
7. Describe your organization’s approach to performing preventative maintenance and site inspections.
a. Describe the frequency at which you typically perform preventative maintenance and site inspections.
b. What tasks does your organization typically perform during preventative maintenance visits?
8. Please detail your organization’s approach to testing DI/RO water systems. What tests does your organization complete and at what frequency?
9. Describe the process by which your organization receives and fulfills requests.
10. Detail any ultra-purification system (polisher) offerings.
11. Describe the service level agreements/key performance indicators that you typically work to achieve.
a. What is your average time to acknowledge a service request?
b. What is your average on-site response time?
c. What is your average on-site response time for emergency service requests?
d. What is your average resolution time?
e. Please describe your system uptime guarantee.
12. Provide details surrounding your quality assurance program.
13. Does your organization maintain an inventory of filters, bulbs, and other commonly needed replacement parts?
14. Should shortages occur, please describe your backup program, including relationships with other manufacturers or vendors to obtain products.
15. How does your organization stay up to date with the latest technologies and best practices in maintaining DI/RO water purification systems?
16. Please describe how your organization keeps detailed records of service history, maintenance completed, and system performance.
17. Provide an example of your organization’s reporting capabilities related to the maintenance and performance of DI/RO Water Purification Systems.
a. Provide examples of service tickets.
b. Provide examples of site-specific service logs.
18. Please describe your approach to decreasing the cost of maintenance and repairs.
19. Describe an overview of the value-added services that you can offer to UA.
20. Please detail your organization’s ability to support the design and installation of replacement DI/RO water purification systems.

## Sections F to H
Implementation | Answer:
1. Describe your distinct implementation and transition plans for UA, including program/project structure, timeline, and plan to minimize disruption of services.
2. If applicable, please describe your approach to taking over work from the incumbent supplier.
Pricing and Billing | Answer:
1. Please refer to the Official Bid Price Excel Workbook attachment to describe the pricing associated with the DI/RO Water Purification System solution.
2. Does your organization offer any early payment discount (e.g., 2% 10 Net 30)?
3. Describe the billing process to UA for services rendered.
4. Describe any benefits offered for longer initial contracted terms.
a. Please describe any benefits offered / impacts to pricing associated with a 5-year term commitment.
5. Please describe any rebates or signing bonuses offered by your organization.
Additional Information (Optional) | Answer:
1. Identify any additional pertinent information that was not considered in the RFP.

## Site Pricing
Official Price Sheet for Exhibit A - Site Pricing
Lists site locations, quantity on hand, part descriptions, and estimated annual exchanges for site inspection and verification of units. Respondents must include proposed part numbers and (per each) part costs & exchange fees.
Building | Location | Type | Part Description | QTY | Frequency | Respondent Part Number (If Applicable) | Single Unit Cost | Exchange Fee (If Applicable)
ADSB | Administrative Services Building | DI | SDI Carbon 1.2cf | 1 | 4 Months
| Humidifier Area |  | SDI Cation 1.2cf | 1 | 4 Months
|  |  | SDI Anion 1.2cf | 1 | 4 Months
AGRI | Agriculture Building | DI | SDI Carbon 1.2cf | 1 | 4 Months
| Room 109 |  | SDI Cation 1.2cf | 1 | 4 Months
|  |  | SDI Anion 1.2cf | 1 | 4 Months
|  |  | SDI Mixbed 1.2cf | 1 | 4 Months
|  |  | Cart Filter 9.75" 5um | 1 | 4 Months
|  |  | Replacement Kit 4 bowl Milli | 1 | 12 Months
|  |  | SDI Mixbed 1.2cf | 1 | Polisher
BELL | Bell Engineering Center | DI | SDI Mixbed .45cf | 1 | 6 Months
| Room 3129 |  | SDI Carbon .25cf | 1 | 6 Months
|  |  | Cart Filter 9.75 1um | 1 | 6 Months
|  |  | SDI Mixbed .45cf | 1 | Polisher
BELL | Bell Engineering Center | DI | SDI Mixbed 1.2cf | 1 | 6 Months
| Proon 2241 |  | SDI Carbon 1.2cf | 1 | 6 Months
|  |  | Cart Filter 9.75 5um | 1 | 6 Months
|  |  | Filter 20" BB | 1 | 12 Months
|  |  | Filter 20" Big | 1 | 12 Months
|  |  | Carb Filter 20" BB | 1 | 12 Months
|  |  | Filter 20" 0.2um Hllw Fbr | 1 | 12 Months
CHBC | Chemistry and Biochemistry Research Building | RO / DI | SDI Mixbed 1.2cf | 1 | 4 Months
| Penthouse |  | SDI Mixbed 1.2cf | 1 | Polisher
|  |  | Filter 20" 5um Code F | 1 | 1 Month
|  |  | Filter 10" 5um BB | 1 | 1 Month
|  |  | Salt Pellet 50# Bag | 5 | 1 Month
|  |  | Readings | 1 | 1 Month
CHEM | Chemistry Building | RO / DI | SDI Mixbed 1.2cf | 1 | 4 Months
| Room 051 |  | SDI Mixbed 1.2cf | 1 | Polisher
|  |  | Filter Cart 9.75 5um | 1 | 1 Month
|  |  | Lamp UV Sterilizer 3084 | 2 | 12 Months
|  |  | Salt Pellet 50# Bag | 15 | 1 Month
|  |  | MMF Performance Data | 1 | 1 Month
|  |  | Filter Performance Data | 1 | 1 Month
|  |  | Softener Performance Data | 1 | 1 Month
|  |  | RO Performance Data | 1 | 1 Month
|  |  | UV Performance Data | 1 | 1 Month
|  |  | Storage Tank Perform Data | 1 | 1 Month
|  |  | DI Performance Data | 1 | 1 Month
CHPN | Champions Hall | RO / DI | SDI Mixbed 3.6cf | 1 | 3 Months
|  |  | Cart 10" 0.2um | 1 | 12 Months
|  |  | Cart 20" 5um BB | 1 | 3 Month
|  |  | UV Lamp 30 in | 2 | 12 Months
|  |  | Salt Pellet 50# Bag | 20 | 1 Month
|  |  | Review Ops Log | 1 | 1 Month
|  |  | Filter Performance Data | 1 | 3 Month
|  |  | Softener Performance Data | 1 | 3 Month
|  |  | RO Performance Data | 1 | 3 Month
|  |  | Function Check PMI | 1 | 3 Month
|  |  | Softener Function Check | 1 | 3 Month
|  |  | DI Maintenance PMI | 1 | 1 Month
DISC | Discovery Hall | DI | SDI Carbon 1.2cf | 1 | 4 Months
| Penthouse |  | SDI Cation 1.2cf | 1 | 4 Months
|  |  | SDI Anion 1.2cf | 1 | 4 Months
|  |  | SDI Mixbed 1.2cf | 1 | 4 Months
|  |  | Filter 9.75" 5um | 1 | 4 Months
|  |  | SDI Mixbed 1.2cf | 1 | Polisher
DISC | Discovery Hall | DI | SDI Mixbed .45cf | 1 | 6 Months
| Room 415 |  | Filter 9.75" 5um | 1 | 6 Months
DISC | Discovery Hall | DI | SDI Mixbed .45cf | 1 | 6 Months
| Room 403 |  | Filter 9.75" 5um | 1 | 6 Months
ENRC | Engineering Research Center | RO / DI | SDI Mixbed 3.6cf | 1 | 3 Months
| Room 364B |  | SDI Mixbed 3.6cf | 1 | Polisher
|  |  | Filter Cart 9.75" 1um | 1 | 3 Months
|  |  | Filter Cart 10" 0.1um | 7 | 12 Months
|  |  | Salt Pellet 50# Bag | 10 | 3 Month
|  |  | Softener Performance Data | 1 | 1 Month
|  |  | RO Performance Data | 1 | 1 Month
|  |  | DI Performance Data | 1 | 1 Month
|  |  | UV Lamp 3084 | 2 | 12 Months
ENRC | Engineering Research Center | RO / DI | SDI Mixbed 3.6cf | 3 | 4 Months
|  |  | SDI Mixbed 3.6cf | 3 | Polisher
|  |  | SDI Mixbed 1.2cf | 1 | 2 Months
|  |  | Filter 10" 5um BB | 1 | 2 Months
|  |  | Filter Cart 30" 0.2um | 7 | 12 Months
|  |  | Salt Pellet 50# Bag | 5 | 2 Months
|  |  | Softener Performance Data | 1 | 1 Month
|  |  | RO Performance Data | 1 | 1 Month
|  |  | UV Lamp 3084 | 4 | 12 Months
FERR | Daniel E. Ferritor Hall | RO / DI | SDI Mixbed 1.2cf | 1 | 1 Month
|  |  | SDI Mixbed 1.2cf | 1 | Polisher
|  |  | Filter Cart 9.75" 5um | 1 | 1 Month
|  |  | Filter Cart 10" 1um | 1 | 6 Months
|  |  | Salt Pellet 50# Bag | 5 | 1 Month
|  |  | Cart Capsule 0.1um Vent | 1 | 12 Months
|  |  | RO Performance Data | 1 | 1 Month
|  |  | Prefilter Maint | 1 | 1 Month
|  |  | Vent Filter Maint | 1 | 12 Months
|  |  | UV Lamp 3084 | 1 | 12 Months
HEAT | Central Utility Plant - Heating Plant | DI | SDI Mixbed 1.2cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
HEAT | Central Utility Plant - Heating Plant | DI | SDI Mixbed .45cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
HPER | Health, Physical Education, and Recreation Building | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room 321T |  | SDI Carbon .25cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
HPER | Health, Physical Education, and Recreation Building | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room 322 |  | SDI Carbon .25cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
|  |  | Ultra Pure DI Pak | 2 | 6 Months
HPER | Health, Physical Education, and Recreation Building | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room 321M |  | SDI Carbon .25cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
LINX | Library Annex | DI | SDI Mixbed 1.2cf | 1 | 6 Months
|  |  | Filter Cart 9.75" 5um | 1 | 6 Months
MULN | David W. Mullins Library | DI | SDI Cation 1.2cf | 1 | 4 Months
| Mechanical Room 115A |  | SDI Anion 1.2cf | 1 | 4 Months
|  |  | SDI Mixbed 1.2cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
NANO | Nanoscale Material Science and Engineering Building | RO / DI | SDI Mixbed 3.6cf | 4 | 4 Months
| Ist Floor RO |  | SDI Mixbed 3.6cf | 4 | Polisher
|  |  | Filter Cart 30" 0.05um | 7 | 12 Months
|  |  | Filter Cart 20" BB 5um | 1 | 1 Month
|  |  | Filter Cart 20" 3um | 7 | 6 Months
|  |  | UV Lamp 30" long | 4 | 12 Months
|  |  | Filter Cart 20" 0.2um Vent | 2 | 12 Months
|  |  | Replace KIT 4 bowl Milli | 1 | 1 Month
GEAR | G. David Gearhart Hall | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room 264 |  | filter Cart 9.75" 5um | 1 | 4 Months
GEAR | G. David Gearhart Hall | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room 265 |  | Filter Cart 9.75" 5um | 1 | 4 Months
GEAR | G. David Gearhart Hall | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room 266 |  | Filter Cart 9.75" 5um | 1 | 4 Months
PHYS | Physics Building | DI | SDI Mixbed 3.6cf | 1 | 4 Months
| Basement |  | Filter Cart 10" 5um BB | 1 | 4 Months
PHYS | Physics Building | DI | SDI Mixbed 1.2cf | 1 | 6 Months
| Room 104 |  | Filter Cart 9.75" 5um | 1 | 6 Months
PHYS | Physics Building | DI | SDI Mixbed .45cf | 1 | 3 Months
| Room 226 |  | Filter Carb Briq 10" 5um | 1 | 3 Months
POHL | John Kirkpatrick Skeeles Poultry Health Laboratory | DI | SDI Mixbed .45cf | 1 | 1 Month
| Bacteriology |  | Filter Carb Briq 10" 5um | 1 | 1 Month
POSC | John W. Tyson Building | RO / DI | SDI Mixbed 1.2cf | 1 | 4 Months
| Basement |  | SDI Mixbed 1.2cf | 1 | Polisher
|  |  | Filter Cart 9.75" 10" 5um | 1 | 1 Month
|  |  | Salt Pellet 50# Bag | 5 | 1 Month
|  |  | RO Performance Data | 1 | 1 Month
|  |  | UV Lamp 3050 | 1 | 12 Months
ROSE | Harry R. Rosen Alternative Pest Control Center | DI | SDI Carbon 1.2cf | 1 | 4 Months
| Alternative Pest |  | SDI Cation 1.2cf | 1 | 4 Months
|  |  | SDI Anion 1.2cf | 1 | 4 Months
|  |  | SDI Mixbed 1.2cf | 1 | 4 Months
SCEN | Science Engineering Hall | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room 504 |  | Filter Cart 9.75" 5um | 1 | 4 Months
PTSC | Plant Sciences Building | DI | SDI Carbon 1.2cf | 1 | 4 Months
| Penthouse |  | SDI Cation 1.2cf | 1 | 4 Months
|  |  | SDI Anion 1.2cf | 1 | 4 Months
|  |  | SDI Mixbed 1.2cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
|  |  | SDI Mixbed 1.2cf | 1 | Polisher
VETD | Veterinary Diagnostic Laboratory | DI | SDI Carbon 1.2cf | 1 | 6 Months
| 2200 West Dean |  | SDI Cation 1.2cf | 1 | 6 Months
|  |  | SDI Anion 1.2cf | 1 | 6 Months
|  |  | SDI Mixbed 1.2cf | 1 | 6 Months
|  |  | Filter Cart 9.75" 5um | 1 | 6 Months
|  |  | SDI Mixbed 1.2cf | 1 | Polisher
sd
AFLS | Agricultural, Food, and Life Sciences Building | DI | SDI Carbon .25cf | 1 | 4 Months
| Room A103 |  | SDI Mixbed .45cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
|  |  | SDI Mixbed .45cf | 1 | Polisher
AFLS | Agricultural, Food, and Life Sciences Building | DI | SDI Carbon .25cf | 1 | 3 Months
| Room C105 |  | SDI Mixbed .45cf | 1 | 3 Months
|  |  | Filter Cart 9.75" 5um | 1 | 3 Months
AFLS | Agricultural, Food, and Life Sciences Building | DI | SDI Carbon .25cf | 1 | 6 Months
| Room C116 |  | SDI Mixbed .45cf | 1 | 6 Months
|  |  | Filter Cart 9.75" 5um | 1 | 6 Months
AFLS | Agricultural, Food, and Life Sciences Building | DI | SDI Mixbed .45cf | 1 | 12 Months
| Room105A |  | Filter Cart 9.75" 5um | 1 | 12 Months
AFLS | Agricultural, Food, and Life Sciences Building | DI | SDI Mixbed .45cf | 1 | 12 Months
| Room105B |  | Filter Cart 9.75" 5um | 1 | 12 Months
AFLS | Agricultural, Food, and Life Sciences Building | DI | SDI Mixbed .45cf | 1 | 12 Months
| Room107 |  | Filter Cart 9.75" 5um | 1 | 12 Months
AGRC | Altheimer Greenhouses | DI | SDI Mixbed 3.6cf | 1 | 4 Months
AGRC | Altheimer Greenhouses | DI | SDI Mixbed 3.6cf | 1 | 4 Months
AGRC | Altheimer Greenhouses | DI | SDI Mixbed 3.6cf | 1 | 6 Months
AGRC | Altheimer Greenhouses | DI | SDI Mixbed 3.6cf | 1 | 4 Months
ALTH | Altheimer Laboratory | DI | SDI Carbon 1.2cf | 1 | 4 Months
|  |  | SDI Cation 1.2cf | 1 | 4 Months
|  |  | SDI Anion 1.2cf | 1 | 4 Months
|  |  | SDI Mixbed 1.2cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
BIOR | Biomass Research Center | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room 132 |  | Filter Cart 9.75" 5um | 1 | 4 Months
BIOR | Biomass Research Center | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room 102 |  | SDI Carbon .25cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
BIOR | Biomass Research Center | DI | SDI Mixbed .45cf | 1 | 2 Months
| Room 117 |  | Filter 10" 20um | 1 | 2 Months
CROP | Crop Science Research Center | DI | SDI Mixbed .45cf | 1 | 3 Months
| Room 158 |  | Filter Cart 9.75" 5um | 1 | 3 Months
CROP | Crop Science Research Center | DI | SDI Mixbed .45cf | 1 | 4 Months
Room 158 Hall
DTAS | Don Tyson Center for Agricultural Sciences | DI | SDI Mixbed 1.2cf | 1 | 6 Months
|  |  | Filter Cart 9.75" 5um | 1 | 6 Months
DTAS | Don Tyson Center for Agricultural Sciences | DI | SDI Mixbed .45cf | 1 | 4 Months
|  |  | SDI Carbon .25cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
DTAS | Don Tyson Center for Agricultural Sciences | DI | SDI Mixbed 1.2cf | 1 | 6 Months
| Room 133 |  | SDI Mixbed 1.2cf | 1 | Polisher
|  |  | Filter Cart 9.75" 5um | 1 | 6 Months
|  |  | Filter 20" BB | 2 | As Needed
|  |  | Filter Cart 10" 0.1um | 1 | As Needed
|  |  | Filter CTR 20" 0.05um UF | 1 | As Needed
DTAS | Don Tyson Center for Agricultural Sciences | DI | SDI Mixbed .45cf | 1 | 6 Months
|  |  | Filter Carbon Briq 10" 5um | 1 | 6 Months
DTAS | Don Tyson Center for Agricultural Sciences | DI | SDI Mixbed .45cf | 1 | 6 Months
|  |  | Filter Carbon Briq 10" 5um | 1 | 6 Months
DTAS | Don Tyson Center for Agricultural Sciences | DI | SDI Mixbed .45cf | 1 | 6 Months
|  |  | Filter Carbon Briq 10" 5um | 1 | 6 Months
DTAS | Don Tyson Center for Agricultural Sciences | DI | SDI Mixbed .45cf | 1 | 6 Months
|  |  | Filter Carbon Briq 10" 5um | 1 | 6 Months
DTAS | Don Tyson Center for Agricultural Sciences | DI | SDI Mixbed .45cf | 1 | 6 Months
|  |  | Filter Carbon Briq 10" 5um | 1 | 6 Months
DTAS | Don Tyson Center for Agricultural Sciences | DI | SDI Mixbed .45cf | 1 | 6 Months
|  |  | Filter Cart 9.75" 5um | 1 | 6 Months
DTAS | Don Tyson Center for Agricultural Sciences | DI | SDI Mixbed .45cf | 1 | 6 Months
|  |  | Filter Carbon Briq 10" 5um | 1 | 6 Months
FARM | FarmHouse | DI | SDI Mixbed .25cf | 1 | 4 Months
| Farm |  | Filter Cart 9.75" 5um | 1 | 4 Months
FDSC | Food Science Building | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room A-5 |  | Filter Cart 9.75" 5um | 1 | 4 Months
FDSC | Food Science Building | DI | SDI Mixbed .45cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
FDSC | Food Science Building | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room C-6 |  | Filter Carbon Briq 10" 5um | 1 | 4 Months
FDSC | Food Science Building | DI | SDI Mixbed 1.2cf | 1 | 6 Months
| Room D-3 |  | Filter Cart 9.75" 5um | 1 | 6 Months
FDSC | Food Science Building | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room N102 |  | SDI Mixbed .45cf | 1 | Polisher
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
FDSC | Food Science Building | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room N104 |  | SDI Mixbed .25cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
FDSC | Food Science Building | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room N110 |  | SDI Mixbed .25cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
FDSC | Food Science Building | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room N111 |  | Filter Cart 9.75" 5um | 1 | 4 Months
0
INRF |  | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room E7 |  | SDI Mixbed .25cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
INRF | Insect Rearing Facility | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room E9 |  | Filter 10" 20um Code F | 1 | 4 Months
INRF | Insect Rearing Facility | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room E11 |  | SDI Mixbed .25cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
POHL | John Kirkpatrick Skeeles Poultry Health Laboratory | DI | SDI Mixbed .45cf | 1 | 4 Months
| Virology |  | SDI Mixbed .25cf | 1 | 4 Months
|  |  | Filter Cart 9.75" 5um | 1 | 4 Months
POSC | John W. Tyson Building | DI | SDI Mixbed .45cf | 1 | 4 Months
| Room L404 |  | Kit 4 bowl Milli | 1 | 12 Months
ROSE | Harry R. Rosen Alternative Pest Control Center | DI | SDI Carbon 1.2cf | 1 | 4 Months
| Alternative Pest |  | SDI Cation 1.2cf | 1 | 4 Months
|  |  | SDI Anion 1.2cf | 1 | 4 Months
|  |  | SDI Mixbed 1.2cf | 1 | 4 Months

## Labor Pricing
Official Price Sheet for Exhibit B - Labor Pricing
Respondent to submit pricing as requested below:
Description | Rate
Standard Hourly Labor Rate (Required Repairs and Maintenance)
Emergency Hourly Labor Rate (Required Repairs and Maintenance)
Hourly Rate for New Installation Services
Travel and related expenses
Other Fee (If Applicable)
Other Fee (If Applicable)
Other Fee (If Applicable)
Other Fee (If Applicable)
Other Fee (If Applicable)

## Parts Pricing
Official Price Sheet for Exhibit C - Parts Pricing
Parts Pricing - Vendor must list a (%) percentage discount from vendor's suggested retail price list that will cover any additional parts or new equipment that is not currently listed at the site locations in Section A. This section will contain 2 categories: 1) consumable parts and 2) new systems (capital equipment). Vendor must supply percentage discount from vendor's most current suggested retail price list and must be able to verify this pricing from invoices or other viable documents. In addition to listing a percentage discount from vendor's suggested retail price list, respondents may choose to provide pricing utilizing the cost + markup fee structure or may propose an alternative fee structure.
Description | Percentage (%) Discount | Insert alternate proposed incentive structure here
Consumable parts percentage (%) discount from vendor's most current suggested retail price list
New system percentage (%) discount from vendor's most current suggested retail price list
Description | Markup %
Consumable parts cost plus markup
New systems parts cost plus markup